[image: image1.jpg]


[image: image2.jpg]




95年公視手持式頻譜儀案規範說明書
1、 通則

1-1 承商須負責進口業務承辦和設備之運送、保險、測試等工作，如需安裝測試，其所需工具及儀表概由承商負責。　

1-2 承商資格如為達到本案所列設備功能完整運作所需者，承商須無條件免費提供相關配件及組裝材料。

1-3 持有台灣地區之合法營利事業登記證，公司執照。

1-4 納稅及無退票紀錄之證明。

2、 設備規格

2-1 Frequency

2-1-1 Frequency Range：100 kHz to 3.0 GHz
2-1-2 Frequency Reference：Aging:(1 ppm/yr, Accuracy:(2 ppm(0~30ºC)
2-1-3 Frequency Span：0Hz, 100 Hz to 2.99 GHz
2-1-4 Frequency Counter：Resolution:10 Hz
2-1-5 Sweep Time：≦1.1 sec(full span)
≦1 msec and ≧20 sec(span=0 Hz)
2-1-6 Resolution bandwidth(-3 dB width)：
1 KHz to 300 KHz in 1、3 steps,(5%
1 MHz,(10%
2-1-7 Video bandwidth(-3 dB)：10 Hz to 1 MHz in 1、3 steps

2-1-8 SSB Phase Noise：<75 dBc/Hz(30 KHz from carrier)
<100 dBc/Hz(100 KHz from carrier)
<120 dBc/Hz(1 MHz from carrier)

2-2 Amplitude

2-2-1 Measurement Range：+20 dBm to -120 dBm

2-2-2 Input Attenuator Range：0 to 30 dB
(selected manually or automatically coupled to the reference level.)
2-2-3 Displayed Average Noise Level(Input terminated, 100 Hz RBW)：
≦-120 dBm, ≧10 MHz and <2.5GHz (preamp on)
≦-115 dBm, ≧2.5GHz (preamp on)

2-2-4 Dynamic Range：≧50 dB, typical
2-2-5 Spurious Responses Input Related：≦-45 dBc
2-2-6 Scale Units：dBm,dBμV,dBmV,μV,mV,V,μW,mW,W
2-2-7 Display Detectors：maximum peak,minimum peak,sample,RMS

2-2-8 Total Level Accuracy：(1.5 dB typical, 10 MHz to 3 GHz
(2 dB typical, <10 MHz
(for input signal levels ≧-60 dBm)

2-2-9 Number of Markers or Delta Markers：6
2-2-10 Marker Functions：peak, minimum,
center=marker frequency,

2-2-11 Marker Displays：normal (level), noise marker,
2-2-12 Trigger：free-running, video, external
2-2-13 Audio Demodulation：AM and FM

2-3 General

2-3-1 Internal Trace Memory：200
2-3-2 Display Traces：active trace and at least 1 memory trace operating at a time

2-3-3 RF Out：Type N, female, 50 ohm

2-3-4 Maximum Input without Damage：+20 dBm, (50 VDC
2-3-5 RF In：Type N, female, 50 ohm

2-3-6 Maximum Input without Damage：+20 dBm, (50 VDC

2-3-7 RF Input VSWR：1.5, typical
2-3-8 Power Supply：External DC Input:meet car providing voltage
External AC Input:100 to 240 Vac, 50 to 60Hz
Internal:NiMH battery, above 1600 mAh
(3 hr operating time for fully-charged battery)
2-3-9 Display Panel：LCD Color
2-3-10 Serial Interface：RS-232, 9 pin D-sub, three wire serial

2-3-11 Weight：≦5 kg(not including accessories)
3、 儀表需求

3-1 採購數量為兩部。

3-2 儀表在環境溫度為0ºC~40ºC及相對溼度（在溫度為35ºC）為85%以內仍可正常運作。

3-3 儀表具有Channel Power Measurement功能、
　　　　VSWR Measurement功能(range:1.0 to 50.0,
　　　　Resolution:0.1dB, frequency range: 25 MHz to 3 GHz)、
　　　　Distance-to-fault Measurement功能
　　　　(distance range: 3 m to 1000 m)。

3-4 儀表具有Power Sensor，可測量+36dBm大小之訊號。

3-5 儀表須提供電腦連線軟體，其系統至少須支援Windows NT/2000/XP，利用其軟體與電腦連線後，可進行測試資料之下載。

3-6 相關測量配件應齊全，每部附有matching pad 50/75ohm、12V car cigarette-lighter adapter、carrying bag、spare battery pack等選項配件各一件，共計兩套，另N(m) to BNC(f) adapter以及N(m) to SMA(f) adapter(0 to 4 GHz)每部儀表各附兩個。

3-7 提供原廠儀表測試報告各兩份。

4、 交貨驗收

4-1 承商須提出設備於四個月內製造完成之證明，且應為未使用之新品。

4-2 儀表交貨時即檢具原廠儀表測試報告，經公視審查通過後辦理驗收。是否驗收合格以提供之原廠儀表測試報告是否符合規格為依據。
4-3 儀表技術手冊operation manual各兩套。

4-4 儀表相關訓練共一日，訓練地點由本會通知，訓練所需教材全數由承商提供。

5、 保固

5-1 驗收完成日起算兩年內，除不可抗拒之天然災害因素，承商應免費修復或換新。

5-2 驗收完成日起算的兩年內，若有firmware版本升級或更新，承商應予免費提供。

5-3 保證八年內繼續供應儀表零件之服務，以利儀表維修或校正。

公視工程部　製


　95.07.21　　v.35





公視手持式頻譜儀案規範說明書














�














































































































PAGE  
5
[image: image1.jpg]                                             　　　　　　　　　　　

